## Slide 1
Organics WG
POM: (umol / kg)
Best PC and PN : High Temp Combustion via Elemental Analyzer
POP: Ash Hydrolysis
Collection: Volume filtered is site dependent
Best: separate bottle dispensed into carboy that can be mixed and sampled to avoid bias against rapidly sinking particles.
Good: subsample bottle – esp in oligotrophic systems
Some sites- prescreen through 200 µm mesh
Filtration: pay attention to filtration rate and avoid overloading filter.
Best- low vacuum (5 in Hg), positive pressure filtration

## Slide 2
POM: PC&PN
Sample prep: when drying sample be careful of oven temperature as drying at 70° C can result in N loss
For Carbon:
Best: run total C and PIC directly to assess POC
Good: acidify (fume) to rid sample of PIC - Validate
Intercalibration:
Best: collect and prepare “in house” references from local system
• Run in every run multiple time in a run
• Swap with other TS for intercalibration
Good: use other references like apple leaf for recovery checks

## Slide 3
DOM (umol /kg):
DOC:
Best: High Temperature Combustion (black art)
Good: Persulfate Oxidation
Acceptable: UV oxidation
TDN:
Best: High Temperature Combustion –Gives TDN must subtract out DIN. Problems when NO3 is too high due to propagation of error
Good: Persulfate Oxidation
Acceptable: UV oxidation
DOP: Best Persulfate Oxidation
Good/ acceptable: UV oxidation - problem with Si interference

## Slide 4
DOM: Filtering is recommended in coastal systems with high POC load. Oligotrophic systems POM load is small and TOM may be appropriate.
Collection: Minimize handling
Best- when filtering best to be conducted directly off Niskin through combusted GF/F in PC cartridge
using silicone tubing…no Tygon!
Combusted Glass or acid leached HDPE bottle
Freeze at -20 up right or acidify to pH 3 and refrigerate
good- subsample and filter in lab. Take special care to minimize contamination!!

## Slide 5
Intercalibration & Consistency between Sites:
There is no certified reference material but Hansell Consensus ref Material are available.
Best/ must: Make 2 – 3 large volume (8 – 10 L) in house References that are calibrated against the Hansell Consensus Material.
These in house references along with blank water should be run every 8- 10 samples to assess performance of the analytical run.
References should not vary by more than 2 – 3 % over the course of an acceptable run
When introducing new “in house” references- analysts need to overlap reference batches.
These references can be swapped among TS for intercalibration of analysis.